R/V Ronald H. Brown METADATA - 2006

Class of Data:  Surface ocean and atmospheric carbon dioxide concentrations

Dataset Identifier:  R/V Ronald H. Brown

One File:  RHB2006

Statement of how to cite dataset:

Ron Brown website: http://www.aoml.noaa.gov/ocd/gcc/rvbrown_data2006.php

These data are made freely available to the public and the scientific community

in the belief that their wide dissemination will lead to greater understanding

and new scientific insights.  The availability of these data does not

constitute publication of the data.  We rely on the ethics and integrity of

the user to assure that AOML receives fair credit for our work.  Please send

manuscripts using this data to AOML for review before they are submitted

for publication so we can insure that the quality and limitations of the data

are accurately represented.

Measurement platform identifier:  NOAA research vessel Ronald H. Brown (R104)

Cruise Information:

The Ron Brown conducted 8 major cruises in the Atlantic and eastern Pacific

Oceans for a total of 11 legs.

Project Information:

The system was operated by personnel from AOML or PMEL (Pacific Marine 	Environmental Laboratory) or by the Ron Brown’s Chief Survey Tech, Jonathan 	Shannahoff.  The work was sponsored by the Underway pCO2 on Ships project of the 	NOAA climate program.

Scientist responsible for technical quality of dataset:

Rik Wanninkhof

NOAA/AOML/Ocean Chemistry Division

4301 Rickenbacker Causeway

Miami, Florida 33149

Rik.Wanninkhof@noaa.gov

Contact person for this dataset:

Bob Castle

NOAA/AOML/Ocean Chemistry Division

4301 Rickenbacker Causeway

Miami, Florida 33149

Robert.Castle@noaa.gov

Timestamp for initial submission of dataset:  11/30/09

Timestamp for the most recent update of dataset:  11/30/09

Timestamp period the dataset refers to: 02/16/2004 – 12/02/2006

Geographic area the dataset refers to:

35 S to 35 N

115 W to 20 W

2006 Cruises:

RB200601 –	Plueddemann Moorings

Charleston, SC to Bridgetown, Barbados

February 16, 2006 to March 3, 2006

Chief Scientist – Al Plueddemann

Operator – Jonathan Shannahoff

RB200602 -	Western Boundary Time Series

Bridgetown, Barbados to Charleston, SC

March 12, 2006 to March 28, 2006

Chief Scientist – Molly Baringer

Operator – Jonathan Shannahoff

RB200603 -	Ocean Explorations

Charleston, SC to San Juan, Puerto Rico

April 11, 2006 to April 30, 2006

Chief Scientist – Peter Wiebe

Operator – Jonathan Shannahoff

RB200604 -	Tsunami Source Potential in the Caribbean

San Juan, Puerto Rico to San Juan, Puerto Rico

May 3, 2006 to May 19, 2006

Chief Scientist – Uri ten Brink

Operator – Jonathan Shannahoff

RB200605A-	AMMA African Monsoon/Sahara Dust Leg A

San Juan, Puerto Rico to Recife, Brazil

May 27, 2006 to June 17, 2006

Chief Scientist – Rick Lumpkin

Operator – Jonathan Shannahoff

RB200605B -	AMMA African Monsoon/Sahara Dust Leg B

Recife, Brazil, to Charleston, SC

June 23, 2006 to July 14, 2006

Chief Scientist – Claudia Schmid

Operator – Jonathan Shannahoff

RB200606A -	Texas Air Quality Study/Gulf of Mexico Atmospheric

Composition and Climate Study Leg A

Charleston, SC to Galveston, TX

July 27, 2006 to August 16, 2006

Chief Scientist – Tim Bates

Operator – Jonathan Shannahoff

RB200606B -	Texas Air Quality Study/Gulf of Mexico Atmospheric

Composition and Climate Study Leg A

Galveston, TX to Galveston, TX

August 22, 2006 to September 11, 2006

Chief Scientist – Tim Bates

Operator – Jonathan Shannahoff

RB200607T -	Stratus Transit

Galveston, TX to Charleston, SC

September 14, 2006 to September 17, 2006

Chief Scientist – N/A

Operator – Jonathan Shannahoff

RB200607 -	Stratus Mooring

Rodman, Panama to Valparaiso, Chile

October 13, 2006 to October 27, 2006

Chief Scientist – Robert Weller

Operator – Jonathan Shannahoff

RB200608 -	TAO

Valparaiso, Chile to Rodman, Panama

November 3, 2006 to December 2, 2006

Chief Scientist – Steve Kunze

Operator – Jonathan Shannahoff

List of variables included in this dataset:

COLUMN	HEADER			EXPLANATION

1.		GROUP/SHIP:			AOML_Brown for all underway data from the Ron Brown.

2.		CRUISE_DESIGNATION:	Cruise ID (e.g., RBYYYYnn where RB = Ron Brown, YYYY

= the four digit year, and nn = the cruise number for

that year).

3.		JD_GMT:			Decimal year day.

4.		DATE_DDMMYYYY:		GMT date.  The date format has been changed to comply

with the IOCCP recommendations.

5.		TIME_HH:MM:SS:		GMT time.

6.		LAT_DEC_DEGREE:		Latitude in decimal degrees (negative values are in 							the southern hemisphere).

7.		LONG_DEC_DEGREE:		Longitude in decimal degrees (negative values are in 							the western hemisphere).

8.		xCO2W_PPM:			Mole fraction of CO2 (dry) in the equilibrator

headspace at equilibrator temperature (Teq) in parts

per million.

9.		xCO2A_PPM:			Mole fraction of CO2 in air in parts per million.

10.		PRES_EQUIL_hPa:		Barometric pressure in the lab in hectopascals (1

hectopascal = 1 millibar).

11.		PRES_SEALEVEL_hPa:	Barometric pressure corrected to sea level from the 							ship’s barometer in hectopascals (1 hectopascal = 1 							millibar).

12.		EqTEMP_C:			Temperature in equilibrator water in degrees

centigade.  Temperature in equilibrator measured with

a calibrated thermistor.

13.		SST(TSG)_C:			Temperature from the ship's thermosalinograph in

degrees centigrade.

14.		SAL(TSG)_PERMIL:		Salinity from the ship's thermosalinograph on the

Practical Salinity Scale.

15.		WATER_FLOW_L/MIN:		Water flow rate through the equilibrator in liters 							per minute.

16.		GAS_FLOW_IR_ML/MIN:	Gas flow through the sample cell of the Licor IR 							analyzer in milliliters per minute.

17.		TEMP_IR_C:			Temperature in the Licor sample cell in degrees 							centigrade.

18.		PRES_IR_hPa:		Barometric pressure in the lab in hectopascals (1

hectopascal = 1 millibar).  The Licor in this system 							does not include a pressure sensor so this field is 							the same as # 10 above.

19.		SHIP_HEADING_TRUE_DEGREE:	Ship’s heading in true degrees from the ship’s 								scientific computing system.

20.		SHIP_SPEED_KNOT:		Ship’s speed in knots from the ship’s scientific 							computing system.

21.		WIND_DIR_REL_DEGREE:	Relative wind direction in degrees from the ship’s 							scientific computing system.

22.		WIND_SPEED_REL_M/S:	Relative wind speed in meters per second from the 							ship’s scientific computing system.

23.		fCO2W@SST_uatm:		Fugacity of CO2 in sea water in microatmospheres

calculated as outlined below.

24.		QC_FLAG_WATER:		Quality control flag for fCO2W@SST measurement.  2 = 							good, 3 = questionable, 4 = bad.

25.		fCO2A_uATM:			Fugacity of CO2 in air in microatmospheres

calculated as outlined below.

24.		QC_FLAG_AIR:		Quality control flag for fCO2A measurement.  2 = 							good, 3 = questionable, 4 = bad.

27.		dfCO2_uATM:			Sea water fCO2 - air fCO2 in microatmospheres.  This

uses the average air value for the current hour.

28.		FLUORO_uG/l:		Measurement from the ship’s Turner 10AU fluorometer 							in micrograms per liter.

29.		WIND_SPEED_TRUE_M/S:	True wind speed in meters per second from the 								ship’s scientific computing system.

30.		WIND_DIR_TRUE_DEGREE:	True wind direction in degrees from the ship’s 								scientific computing system.

31.		AIR_TEMP_C:			Outside air temperature from the ship’s scientific 							computing system.

The following fields have been QC'ed by the CO2 group:

GROUP/SHIP

CRUISE_DESIGNATION

JD_GMT

DATE_DDMMYYYY

TIME_HH:MM:SS

LAT_DEC_DEGREE

LONG_DEC_DEGREE

xCO2W_PPM

xCO2A_PPM

EqTEMP_C

PRES_EQUIL_hPa

WATER_FLOW_L/MIN

GAS_FLOW_L/MIN

TEMP_IR_C

PRES_IR_hPa

fCO2W@SST_uatm

fCO2A_uATM

dfCO2_uatm

The following fields are from the ship's onboard systems and the quality of this

data cannot be verified:

SST(TSG)_C

Sal(TSG)_Permil

PRES_SEALEVEL_hPa

SHIP_HEADING_TRUE_DEGREE

SHIP_SPEED_KNOT

WIND_DIR_REL_DEGREE

WIND_SPEED_REL_M/S

FLUORO_uG/l

WIND_SPEED_TRUE_M/S

WIND_DIR_TRUE_DEGREE

AIR_TEMP_C

Narrative description of system design:

CO2 ANALYTICAL SYSTEM:

The concentration of carbon dioxide (CO2) in surface ocean water is determined

by measuring the concentration of CO2 in gas that is in contact with the water.

Surface water is pumped ~ 100 m through 7/8" Teflon tubing from an inlet

in the ship's bow to the equilibration chamber.  Water comes from the bow

intake ~4.2 m below the water line and the TSG is located close to the inlet.

When the SST is below about 20 oC, friction in the pipes and from the pump cause

heating and the Teq is higher than SST.  When the SST is higher than about 25 oC,

the ship’s air conditioning cools the water and the Teq is lower than SST.

The equilibration chamber has an enclosed volume of gas, or headspace, and a pool

of seawater that continuously overflows to a drain.  As the water flows through the

chamber, the dissolved gases (like CO2) partition between the water and the

headspace.  At equilibrium, the ratio of CO2 in the water and in the headspace is

influenced most by temperature, and that relationship is known.  By measuring

the concentration of CO2 in the headspace and the temperature in the chamber,

the partial pressure (or fugacity) of CO2 in the surface water can be calculated.

INSTRUMENT DESCRIPTION

The general principle of instrumental design can be found in Wanninkhof and Thoning

(1993), Ho et al. (1995), and Feely et al. (1999).  The concentration of CO2 in the

headspace gas is measured using the adsorption of infrared (IR) radiation, which

results from changes in the rotational and vibrational energy state of the CO2

molecule.  The LI-COR detector passes IR radiation through two 6" cells.  The

reference cell is flushed with a gas of known CO2 concentration.  The sample cell

is flushed with the headspace gas.  A vacuum-sealed, heated filament is the

broadband IR source.  The IR radiation alternates between the two cells via a

chopping shutter disc. An optical filter selects an adsorption band specific

for CO2 (4.26 micron) to reach the detector.  The solid state (lead selenide)

detector is kept at -12 degrees °C for excellent stability and low signal

noise (less than 0.2 ppm).

Several steps are taken to reduce interferences and to increase the accuracy

of the measurements.  After the equilibration chamber, the headspace travels

through a drying trap to remove water vapor.  During each analysis, the

headspace gas is compared to a reference gas of known concentration.  To

improve the accuracy of the measurements, three different gaseous standards

for CO2 are analyzed once an hour instead of the headspace gas.

Analyzer:  LI-COR 6251 (analog output) infrared (IR) analyzer.

Method of Analysis:  Differential analyses relative to the low standard.  Measures

dried equilibrator headspace gas.  Gas flow is stopped prior to IR readings.

Drying Method:  The equilibrator headspace sample gas first goes through a glass

condenser cooled to ~ 5 oC.  The sample and standard gases pass through a short

column of magnesium perchlorate before reaching the analyzer.

Equilibrator (setup, size, flows):  The equilibrator is based on a design by R.

Weiss and was fabricated from a plexiglass housing with ~8 L water reservoir and

~16 L gaseous headspace.  Water flow rate is ~11 L/min.  Headspace recirculation

rate is ~200 ml/min.

Additional sensors:

The 10-cm thermistor used to electronically log the temperature was mounted in the

bottom of the equilibrator.  It was calibrated annually against a Guildline model

9540 digital platinum resistance thermometer with a NIST traceable probe, or a

Hart Scientific 1560 Black Stack module with platinum resistance NIST traceable

thermistor.  Based on reproducibility of the annual calibrations, the temperatures

are believed accurate to 0.02 ˚C

The barometric pressure was measured in the lab next to the equilibrator with a

Setra model 370 electronic barometer with an accuracy of ± 0.2 hPa.  Periodic

comparison of barometers gave readings within ± 0.5 hPa several. The equilibrator

had two 0.5-cm ID vents to the laboratory and thus equilibrator headspace pressure

was assumed to be laboratory pressure.

A YSI model 600 R thermosalinograph with temperature, salinity and dissolved

oxygen probe was mounted in the sink next to the equilibrator for diagnostic

purposes.  Temperature from this unit had a precision of 0.05 ˚C but an offset

of 0.2 ˚C.

A Seabird SBE 21 thermosalinograph was mounted in a seachest chamber 4 m from the

intake at nominally 5-m depth. The unit was calibrated annually and provided SST to

better than 0.02 C and salinity generally to 0.1 or better.

The dissolved oxygen measurements are not reported in the final data file.

Narrative statement identifying measurement method for each required parameter:

CALCULATIONS:

The mixing ratios of ambient air and equilibrated headspace air are calculated

by fitting a second-order polynomial through the hourly averaged millivolt

response of the detector versus mixing ratios of the standards.  Mixing ratios

of dried equilibrated headspace and air are converted to fugacity of CO2 in

surface seawater and water saturated air in order to determine the fCO2.

For ambient air and equilibrator headspace, the fCO2a (or fCO2eq) is calculated

assuming 100% water vapor content:

fCO2eq = xCO2eq(P-pH2O)exp(B11+2*d12)P/RT

where fCO2eq is the fugacity in the equilibrator, pH2O is the water vapor

pressure at the sea surface temperature, P is the atmospheric pressure (in atm),

T is the SST or equilibrator temperature (in K) and R is the ideal gas constant

(82.057 cm^3·atm·deg^-1·mol^-1).  The exponential term is the fugacity correction

where B11 is the second virial coefficient of pure CO2

B11 = -1636.75 + 12.0408T - 0.032795T^2 + 3.16528E-5 T^3

and d12 = 57.7 - 0.118 T is the correction for an air-CO2 mixture in units of

cm^3·mol^-1 (Weiss, 1974).

The calculation for the fugacity at SST involves a temperature correction term

for the increase of fCO2 due to heating of the water from passing through the

pump and through 5 cm ID PVC tubing within the ship.  The empirical temperature

correction from equilibrator temperature to SST is:

fCO2(SST) = fCO2(eq) /

Exp ((Teq-SST) * [0.03107 – 2.7851E-4 * Teq – 1.8391E-3 * ln(fco2eq * 1.0E-6)])

where SST is sea surface temperature and Teq is the equilibrator temperature in

degrees °C.

Sampling Cycle:

The system runs on an hourly cycle during which 3 standard gases, 3 air

samples from the bow tower and 8 surface water samples (from the

equilibrator head space) are analyzed on the following schedule:

Mins. after hour		Sample

4 				Low Standard

8 				Mid Standard

12 				High Standard

16.5				Water

21 				Water

25.5				Water

30				Water

34				Air

38				Air

42				Air

46.5				Water

51				Water

55.5				Water

60				Water

NOTES ON DATA:

Columns have a default value of –999.99 in case of instrument malfunction,

erroneous readings or missing data.  Furthermore, if a suspicious xCO2 value,

pressure or temperature value is encountered, the fCO2 is not calculated.

Analytical Instrument Manufacturer/Model:

The Ron Brown system (version 2.6) was built by Craig Neill in 1999.  The analyzer 	is a LI-COR 6251 (analog output) infrared analyzer.

Standard Gases and Reference Gas:  The three standard gases came from CMDL

in Boulder and are directly traceable to the WMO scale.  While individual data

points above the high standard gas concentration or below the low standard gas

concentration may not be accurate, the general trends should be indicative of the

seawater chemistry.

Description of any additional environmental control:

The system is located in the Hydro Lab of the Ron Brown.  The room is

air-conditioned with little temperature fluctuation.

Resolution of measurement:

The resolution of the instrument is better than 0.1 ppm.

Estimated overall uncertainty of measurement:

The xCO2eq measurements are believed accurate to 0.1 ppm.  The

fCO2@SST measurements are believed to be precise to 0.2 ppm.

List of calibration gases used:

The standards used during the 2006 field season were:

STANDARD		TANK #	CONCENTRATION	VENDOR

Low			CA05760	303.23		ESRL

Mid			CA05398	370.90		ESRL

High			CA05344	411.42		ESRL

High			CC06352	514.29		ESRL

Traceability to an internationally recognized scale (including date/place of last calibration made):

All standards are obtained from NOAA/CMDL, now called the Global

Monitoring Division of the Earth System Research Laboratory and are directly

traceable to WMO scale.

Uncertainty of assigned value of each calibration gas:

The uncertainty based on pre and post cruise calibrations is less

than 0.05 ppm.

Pressure/Temperature/Salinity:

For information about the ship’s thermosalinograph or other sensors, contact Chief 	Survey Tech Jonathan Shannahoff at jonathan.shannahoff@noaa.gov.

Units:

All xCO2 values are reported in parts per million (ppm) and fCO2 values

are reported in microatmospheres (uatm) assuming 100% humidity at the

equilibrator temperature.

Bibliography:

DOE (1994). Handbook of methods for the analysis of the various parameters of the

carbon dioxide system in sea water; version 2. DOE.

Feely, R. A., R. Wanninkhof, H. B. Milburn, C. E. Cosca, M. Stapp and P. P. Murphy

(1998). A new automated underway system for making high precision pCO2

measurements onboard research ships. Analytica Chim. Acta 377: 185-191.

Ho, D. T., R. Wanninkhof, J. Masters, R. A. Feely and C. E. Cosca (1997).

Measurement of underway fCO2 in the Eastern Equatorial Pacific on NOAA

ships BALDRIGE and DISCOVERER, NOAA data report ERL AOML-30, 52 pp.,

NTIS Springfield.

Wanninkhof, R. and K. Thoning (1993). Measurement of fugacity of CO2 in surface

water using continuous and discrete sampling methods. Mar. Chem. 44(2-4):

189-205.

Weiss, R. F. (1970). The solubility of nitrogen, oxygen and argon in water and

seawater. Deep-Sea Research 17: 721-735.

Weiss, R. F. (1974). Carbon dioxide in water and seawater: the solubility of

a non-ideal gas. Mar. Chem. 2: 203-215.

Weiss, R. F., R. A. Jahnke and C. D. Keeling (1982). Seasonal effects of

temperature and salinity on the partial pressure of CO2 in seawater.

Nature 300: 511-513.

Comments related to all 2006 data:

1.	xCO2 values outside the range of the standard gases (i.e. below the low standard or	above the high standard) are not as accurate as values within the range.  However,	the general trends should be indicative of the seawater chemistry.

2.	The standard gases for the first 9 legs (RB200601-RB200607T) were 303.23 ppm, 	370.90 ppm, and 411.42 ppm.  The standards for the last 2 legs (RB200607 and	RB200608) were 303.23 ppm, 370.90 ppm, and 514.29 ppm.

Comments related to the individual legs:

RB200601:	1. SST and salinity data from the ship's TSG were bad for the first 6 				hours of the cruise.  They have been replaced with corrected data from 				the Seabird Micro TSG near the underway system and the thermistor 				inside the equilibrator (EqT).  SST was replaced with EqT - 0.16 and 				salinity was replaced with Seabird salinity - 0.02.

2. There was low gas flow in the water phase from 2/18 at 1801 to 2/20 				at 1501.  Most of the water values in this period have been removed.

RB200602:	1. The salinity reading from the ship's TSG was bad from the beginning 				of the cruise until March 17 at 2049.  It was replaced with values 				based on the Seabird TSG near the underway system (Seabird salinity - 				0.106).

2. From March 16 at 1523 to March 17 at 0107 a spike occurred in the 				water xCO2 values.  While the cause of this spike could not be exactly 				determined, the values were removed because it was suspected that they 				were not correct.

3. There were 4 approximately one day long periods where the Licor cell 			temperature was very noisy.  Since the cell temperature is not used in 				calculating xCO2, the values were not removed, but it could be 					indicative of problems developing with the Licor.

RB200603:	1. No data was recorded from April 28 at 1209 to April 29 at 2111.

2. There were 3 periods where the Licor cell temperature was very 				noisy.  Since the cell temperature is not used in calculating xCO2, the 			values were not removed, but it could be indicative of problems 					developing with the Licor.

RB200604:	1. The Licor cell temperature was very noisy.  Since the cell 					temperature is not used in calculating xCO2, the values were not 					removed, but it could be indicative of problems developing with the 				Licor.

RB200605A:	1. The Licor cell temperature was noisy.  Since the cell temperature is 			not used in calculating xCO2, the values were not removed, but it could 			be indicative of problems developing with the Licor.

RB200605B:	1. There were three data dropouts:  on July 6 at 0733 for 3 hours, on 				July 8 at 0120 for 11.5 hours, and on July 12 at 1015 for 29 hours.

2. On July 11 at 0200 there was an abrubt drop in air values from about 			382 to about 377.5 ppm and the values remain constant for about 1.25 				days when there is a data dropout.  After the dropout, the air values 				return to the 382 ppm level for the remainder of the cruise.  A check 				of the air values from the Carbon Cycle Greenhouse Gas group at NOAA's 				Earth System Research Laboratory ( http://www.cmdl.noaa.gov/ccgg/ ) 				shows that the values recorded at the nearest station (St. David's 				Head, Bermuda) for that time were also in the high 370's.

RB200606A:	1. There were seven data dropouts of varying length.  Due to the 					frequent system down time, only about 49% of the total possible data 				was actually measured.

2. There were numerous high air values (above the typical open ocean 				value of ~380 ppm.  Where these could not be attributed to 						contamination by the ship's stack gases, they have been left in and are 			assumed to be caused by shore-based sources.

RB200606B:	1. There were eight data dropouts of varying length.  Due to the 					frequent system down time, only about 46% of the total possible data 				was actually measured.

2. There were numerous high air values (above the typical open ocean 				value of ~380 ppm.  Where these could not be attributed to 						contamination by the ship's stack gases, they have been left in and are 			assumed to be caused by shore-based sources.

RB200607T:	1. There were several high air values, mostly at the beginning and end 				of the cruise.  These are believed to be caused by shore-based sources 				and have been left in the data file.

RB200607:	1. The salinity values from the ship's thermosalinograph were bad and 				were replaced with values from the Seabird Micro TSG at the seawater 				input to the pCO2 system.

RB200606:	No problems of note.