Cruise:  EX0520E

Ship:  M/V Explorer of the Seas

Dates:  May 15 - 22, 2005

Chief Scientist:  not applicable

Equipment:  Surface samples collected.

Total number of stations:  VOS Underway Cruise

Sample Collection

The discrete samples were collected by Heike Lueger from a tap at the sink next to the underway pCO2 instrument.  The underway pCO2 instrument was located in the Ocean Lab, while the TSG was located in the bow thruster space.  The sea water takes an estimated 1.5 minutes to travel from the bow thruster space to the Ocean Lab.  The TSG temperature is believed to be no more than 0.15 degrees C warmer than in-situ SST.

The date and time listed in the data file are UTC when each sample bottle was collected.

DIC:

Sample_ID#:  1 - 20

10 locations, 20 samples each 500-ml, 10 sets of duplicate samples

PI:  Dr. Rik Wanninkhof

Analyzed by:  Esa Peltola

TAlk:

The samples analyzed for DIC were later analyzed for TAlk.

PI:  Dr. Frank Millero

Analyzed by:  Fen Huang

Sample Analysis

DIC:

Analysis date: August 2, 2005

Coulometers used: AOML1 and AOML2

Blank range: 19.8-28.2 counts/min

CRM # used and assigned value (include both DIC and salinity): Batch 59, c: 2007.1 umol/kg,S:33.316

CRM value measured:  AOML 1: offset 3.9 umol/kg ( 2011.0 umol/kg)

AOML-2: offset 0.8 umol/kg ( 2007.9 umol/kg)

Average run time, minimum run time, maximum run time:  12 min, 9 min, 19 min

Reproducibility: (# samples and average difference): 10 sets of duplicate samples, average difference 1.1 umol/kg

CRM, salinity and HgCl2 correction applied: yes

Remarks-

The first CRM of each cell was used for a CRM correction.

Comments

The GPS transducer and the thermosalinographs (TSG) were logged by a computer system installed by NOAA and the University of Miami.  These instruments were maintained by a technician from the University of Miami’s Marine Technology Group (http://www.marinetechnologygroup.org/) sailing on the Explorer of the Seas full-time.  The data from the TSGs, Seabird SBE-21 and SBE-45, are archived by the Ship of Opportunity Program at AOML (http://www.aoml.noaa.gov/phod/tsg/soop/index.php).

The latitude, longitude, temperature and salinity reported with the DIC and TAlk measurements were taken from the raw TSG data file.  The merging of the discrete measurements with the TSG data was done on the basis of date and time.  The TSG values are provided for reference; no post-cruise assurance of accuracy has been done to this data.

The Sample_ID is the sample bottle number for the discrete samples.