Cruise:  EX0645S

Ship:  M/V Explorer of the Seas

Dates:  November 4 - 12, 2006

Chief Scientist:  not applicable

Equipment:  Surface samples collected.

Total number of stations:  VOS Underway Cruise

Sample Collection

The discrete samples were collected by Kevin Sullivan from a tap at the sink next to the underway pCO2 instrument.  The underway pCO2 instrument was located in the Ocean Lab, while the TSG was located in the bow thruster space.  The sea water takes an estimated 1.5 minutes to travel from the bow thruster space to the Ocean Lab.  The TSG temperature is believed to be no more than 0.15 degrees C warmer than in-situ SST.

The date and time listed in the data file are UTC when each sample bottle was collected.

DIC:

Sample_ID#:  1 - 20

10 locations, 20 samples each 500-ml, 10 sets of duplicate samples

PI:  Dr. Rik Wanninkhof

Analyzed by:  Esa Peltola

Sample Analysis

DIC:

Analysis date: December 4, 2006

Coulometers used: AOML1 and AOML2

Blank range: 13.4-28.6 counts/min

CRM # used and assigned value (include both DIC and salinity): Batch 59, c: 2007.1 umol/kg,S: 33.316

CRM value measured:  AOML 1: offset 10.6 umol/kg (2017.7 umol/kg)

AOML-2: offset 10.1 umol/kg (2017.2 umol/kg)

Average run time, minimum run time, maximum run time:  10 min, 8 min, 17 min

Reproducibility: (# samples and average difference): 10 sets of duplicate samples, average difference 1.6 umol/kg

CRM, salinity and HgCl2 correction applied: Salinity correction was applied for the AOML1 samples, the AOML2 samples were calculated using TSG salinity; CRM and HgCl2 volume correction was applied

Remarks-

The volume correction was applied due to added HgCl2 (Measured DIC*1.00037).

The first CRM of each cell was used for a CRM correction.

There was a good agreement between the duplicate samples All 20 samples were sent to Millero's lab for alkalinity analysis

Comments

The GPS transducer and the thermosalinographs (TSG) were logged by a computer system installed by NOAA and the University of Miami.  These instruments were maintained by a technician from the University of Miami’s Marine Technology Group (http://www.marinetechnologygroup.org/) sailing on the Explorer of the Seas full-time.  The data from the TSGs, Seabird SBE-21 and SBE-45, are archived by the Ship of Opportunity Program at AOML (http://www.aoml.noaa.gov/phod/tsg/soop/index.php).

The latitude, longitude, temperature and salinity reported with the DIC measurements were taken from the raw TSG data file.  The merging of the discrete measurements with the TSG data was done on the basis of date and time.  The TSG values are provided for reference; no post-cruise assurance of accuracy has been done to this data.

The Sample_ID is the sample bottle number for the discrete samples.