## Slide 1: Update on the RD41 measurements
NCAR |
EOL
Holger Vömel, T. Hock, M. Goodstein, K. Young, I. Suhr NCAR EOL
J. Smith
NOAA/AOC
20th AVAPS Users Group Meeting 25 April 2018
Notes: Acknowledge NOAA AOC for their effort

## Slide 2: New Vaisala RS41 PTU Module
New PTU Sensors module based on RS41 radiosonde  RD41 dropsonde
New resistive platinum temperature sensor and support structure
One RH sensor
  Humidity sensor has a built-in heater & temperature sensor
Same pressure sensor technology
RD41
RD94
RH2
RH1
Temp
RH
Temp
Notes: the humidity sensor has a built-in temperature sensor which continuously monitors the temperature of the humidity sensor. This temperature is factored in humidity measurements to ensure accuracy in all conditions.

## Slide 3
Temperature

## Slide 4: Faster temperature measurements
RD41 temperature time constant about 4 times faster
Notes: Time lag correction makes a difference of about 0.2 K near the surface to about 0.7 K in upper troposphere

## Slide 5: Difference between RD41 and RD94 temperature
Aspen time lag corrected

## Slide 6: RD41 Temperature Repeatability
17 Sondes

## Slide 7: Comparison to Buoy

## Slide 8
Relative humidity

## Slide 9: RD94 and RD41 relative humidity

## Slide 10: RD41 relative humidity variability
17 Sondes

## Slide 11: NOAA RD41/RD94 intercomparison October 2017

## Slide 12: NOAA RD41/RD94 intercomparison October 2017

## Slide 13: Comparison to Buoy

## Slide 14
Pressure

## Slide 15: NOAA RD41/RD94 intercomparison October 2017

## Slide 16: RD41 pressure variability
17 Sondes

## Slide 17: Comparison to Buoy

## Slide 18
Winds

## Slide 19: RD41 wind variability
17 Sondes

## Slide 20: Near surface vertical fall velocity

## Slide 21: Summary
NCAR dropsonde PTU sensor module updated: RD41
Faster temperature measurement Systematic error to RD94 < 0.2 K Systematic error to buoy < -0.5 K Repeatability < 0.3 K
Better relative humidity measurement RD94 dry bias eliminated Systematic error to RD94 ~ 3% (due to high range) Systematic error to buoy < -3% Repeatability ~ 2% (in lowest km)
No change in pressure measurements Better offset correction Repeatability < 1 hPa Systematic error to buoy < -0.5 hPa
Update in GPS eliminated near surface wind spikes Repeatability ~ 1 – 2 m/s