Flight ID: 070925I

Sensor or system						Variable Name

INE								ine1

Accelerometer							acci1X

Temperature Probe						tt1 (Rosemount FWD)

Dew Point Probe						tdm2X (Edgetech Vigilant)

Altitude (for vertical wind)					ra159 (GPS Aided Inertial Altitiude)

Static Pressure							psf (Rosemount Fuselage)

Dynamic Pressure						pqf1 (Rosemount Fuselage)

Time Source							Micro 99

Local Met Data:		Takeoff (1358Z)		Landing (1925Z)

Aircraft Static Pressure		1016.8 	 mb		1007.2	 mb

Tower Pressure			1016.7	 mb		1013.0	 mb

Notes

There are no data gaps.

From 143112z-143210z the new happy box was modified resulting in a plateau of the data that goes through the happy box on the way to the data tape. The resulting plateaus were not removed, but are found in the following variables: acci1, acci2, dai1, dai2, gsi1, gsi2, gsui1, gsui2, gsvi1, gsvi2, hdgi1, hdgi2, lati1, lati2, loni1, loni2, ptchi1, ptchi2, rolli1, rolli2, ra232, alti2, vsi1, vsi2.

There are spikes in the accelerometer data at: 143210z-143213z and 161526z-161528z. These spikes were removed from the acci1 variable and saved into acci1X to be used in meteorological calculations.

There are spikes in the ashtech lat and lon data at: 191455z-191544z, 182357z-182400z, 182942z-182945z. These spikes were not removed from the data.

There was a spike in the tdm2 data at 142336z-145212z where tdm2 was balancing. I removed this spike and substituted data from tdm1.

There were periods where the dewpoint temperature exceeded the ambient temperature resulting in RH values > 100%. This was likely due to heavy rain, a wet-bulb effect on the total temperature sensor, and/or an artificial warming of the dewpoint sensor as it tried to burn off excess moisture. Corrections were not made to the data.

Aircraft INE positions were re-navigated with respect to GPS

Radar altitude from the HG9550 radar altimeter is in the RA-232 variable slot. GPS aided INE1 altitude is in the RA-159 variable slot. The RA-232 and RA-159 altimeters are no longer installed on the aircraft.

The HG9550 radar altimeter data has a consistent 1% positive bias above 1,500 feet that was discovered after flight. The correction formula is:

(((reported altitude -  1,500) * 0.988) +1,500) = corrected altitude

This formula should only be needed for this hurricane season’s flight data as the problem should be corrected automatically during next season’s flights.

There is no King Liquid Water Content (KLWC) sensor installed.

There is no Maycomm Water Vapor Sensing System (WVSSII) installed

There is no TT3 installed.

All other instruments appeared to work optimally.