N42RF ERROR SUMMARY

H. GUSTAV NEAR WESTERN CUBA

Flight ID: 080830H

Sensor or system						Number or Name

Inertial  							INE1

Accelerometer						ACC1

Temperature Probe				                      TT1

Dew Point Probe						TDM2X (EdgeTech)

Static Pressure						PSF

Dynamic Pressure						PQF1

Vert. Wind						           RA159X

Constants File                                                                  n42_hur08v3.adc

Project Directory						/acdata/2008/hur08

Notes:

There were no data gaps.  However beginning with the H080829 H. Gustav flight and continuing through the H080912 H. Ike flight, there was observed a 3 to 4 second lag between all INE1 and INE2 outputs. INE1was selected since it best corresponded with the analog output.

There were no liquid water sensors available for the flight.

The APN-159 (RA159X) radar altimeter output had some time periods of erroneous values. The RA159 data glitches were fixed using statistical techniques during the following time periods: 212900Z – 213100Z, 010000Z – 010200Z, and

011300Z - 011500Z.

Measured dewpoint temperature #2 (EdgeTech)...TDM2...exhibited three  time periods of erroneous data during the flight at 210358Z – 210428Z,000115Z – 000512Z and

005906Z - 010112Z.

The erroneous data was removed by substituting TDM3(Maycomm/TDL) output,

TDM2 = TDM3

All other instruments worked optimally during the flight.

Nine (9) GPS dropsondes were deployed with all being good. Nine (9) tempdrop messages were sent to NHC.

Takeoff		Landing

2018Z		0153Z

Aircraft Static Pressure                   1009.5mb	1010.8mb

Corrected Tower Pressure               1009.7mb	           1011.2mb

Flight Director:	A. Barry Damiano       (813) 828-3310 ext. 3073

# Text boxes

1