N43RF ERROR SUMMARY

HURRICANE 2009

HURRICANE BILL

Flight ID: 090820I2

Sensor or system						Number or Name

INE	(for wind derivation)				INE1

Accelerometer						ACC1

Temperature Probe				                     TT1

Dew Point Probe						TDM2X (EDGETECH)

Static Pressure						PSF

Dynamic Pressure						PQF1

Vert. Wind						          ALTI1X

Constants File                                                           n43_hur09v2.adc

Project Directory						/acdata/2009/hur09/P3

Notes:

There were 6 instances after takeoff of data gaps in the netCDF file:

00:52:01 – 00:52:11Z: no subsequent disruption resulted from this gap

01:09:11 – 01:09:19Z: during this interval, there were significant errors caused by the interface between the inertial navigation systems and the RAMS data collection system that generates High Density Observations in real time. As a result, the ASDL SATCOM system sent out greatly exaggerated flight level and SFMR wind speeds for HDOBS 01:09:00 and 01:09:30 Z. However, since in the creation of the netCDF file from the aircraft tape this nine second gap appears as missing data, no such errors are included in the research data. Fortunately also, no disruption of subsequent data occurred immediately following this gap in the netCDF file.

There were four more instances of one second gaps:

02:15:51Z

02:38:41Z

02:40:31Z

02:41:01Z

In each of these instances, there were significant spikes in the subsequent second resulting from errors in sampling inertial data output. This necessitated that the following variables be patched using statistical means: alti1, lati1, loni1. Two more: gsui1 and gsvi1 had very small data deviations not requiring patching.

Dewpoint sensor #2...TDM2 (EdgeTech) had several occurrences where its' value was greater than ambient temperature thus producing humidity values above 100%. For most of these occurrences TDM2 output was not modified. There were two instances where TDM2 had to be patched using statistical methods (data smoothing to remove anomalous spikes with weighting factors between 0.10 and 0.30) to avoid having RH values well in excess of 120% in the final RXC file: 23:20:00-23:20:33Z and 02:44:06 – 02:45:08Z.

The tail Doppler radar was not available for this mission. All other flight level instruments worked optimally during the flight. There was an abnormally long period of ground time between when the data system began collecting measurements and takeoff due to a maintenance delay (over 90 minutes from 1940z to 2122z).

Twenty dropsondes (20) were deployed during the mission. Two were bad (the 0022Z drop at the NW corner of the pattern was a fast fall and the 0143Z drop at the 060 degree radial out from eye along the arc pattern had no launch detect).  No AXBT were deployed.

The RINU1550 GPS altitude output was used for extrapolating sea level pressure from flight altitude. An eye center dropsonde was conducted at 23:58:28Z. The extrapolated sea level pressure from flight altitude, 700 mb or 10,000 feet, at the time of this release was 940.3 mb with a splash sonde pressure (according to ASPEN, Editsonde was not run on this drop) of 943 mb (with 8 knots of wind at the surface).

SPECIAL NOTE!!!     The variable names dpj wgs, dpj_was and dpj_wz in the netCDF file represent vertical ground, vertical air and vertical wind speeds, respectively, computed using Dave Jorgensen’s vertical wind algorithm.  It is recommended that these values be used for vertical wind analysis.

Takeoff		Landing

2122Z                    0512Z

Aircraft Static Pressure                1004.5mb	  1007.0mb

Corrected Tower Pressure           1003.7mb	  1007.8mb

Flight Directors: Richard Henning and Jack Parrish (813) 828-3310 ext. 3086

# Text boxes

1