Flight ID: 20110823N1

Sensor or system						Number or Name

Accelerometer							AccZI.1

Altitude 							AltGPS.3

Attack Angle		AA.2

Dew Point Probe						TdM.1X

Dynamic Pressure						PQF.2

Geopotential Altitude						ALTGA.1 (AltGPS.3)

Inertial Selected						INE 1

Static Pressure							PSF.2

Slip Angle		SA.2

Temperature Probe						TTM.1

True Air Speed		TASF.1

Constants File		49cal111.doc

Flight Directory		acdata/MET/2011/20110223N1

Local Met Data:			Takeoff (0520Z)		Landing (1221Z)

Aircraft Static Pressure		1013.5 mb			1011.9 mb

Tower Pressure (corrected)		1014.1 mb			1012.7 mb

Notes:

TTM.3 had an oscillation throughout the flight with the same period. The amplitude of the oscillation decreased as altitude increased resulting in about 0.8°C amplitude from 40-45kft.  TTM.1 is default.

PQBetaF.1, Dynamic Slip Pressure 1, spikes downward from 115326Z-115645Z, causing a spike in SA.1 which caused a spike in UWZ and WS.  Due to this spike PQBetaF.2/SA.2 was selected as default.

Dew point values intermittently exceeded ambient temperature values from 0811Z – 0926Z during portions of flight where the aircraft was in precipitation, causing RH values greater than 100%.  TdM.1 performed better overall maintaining a slightly colder temperature at altitude and compared well with ground station dew point.  However, it had a dropout on descent from 120638Z-121809Z, during this time TdM.2 was directly substituted in for TdM.1, TdM.2 = TdM.1.  TdM.1X is default.

All other instruments worked properly.

26 dropsondes launched (AVAPS II): 23 good, 3 terminated early, 26 sent.

All 26 good drops successfully transmitted to be ingested into the 23/12Z model.