Flight ID: 20120823N1

Sensor or system						Number or Name

Accelerometer							AccZI.1

Altitude 							AltGPS.3x

Attack Angle		PDAlpha.2

Dew Point Probe						TDM.1

Dynamic Pressure						PQM.2

Geopotential Altitude						ALTGA.1 (AltGPS.3x)

Inertial Selected						INE.1

Static Pressure							PSM.2

Slip Angle		PDBeta.2

Temperature Probe						TTM.1

Constants File		49cal102

Flight Directory		acdata/MET/2012/20120823N1

Local Met Data:			Takeoff (1731Z)		Landing (0140Z)

Aircraft Static Pressure		1018.0mb			1015.5mb

Tower Pressure (corrected)		1019.2mb			1017.5mb

Notes:

At 00z there was a system wide data gap that appeared from 00:00:06z to 00:00:53z.

TDM.1 and TDM.2 are both not rated for use under -50 deg C.  Cannot be considered reliable for dewpoints colder than -50C.

ALTPA was using an incorrect equation within AAMPs for the first portion of the flight causing a large offset from 17:49.52z until 19:49:49 when the correct equation was implemented.

There were 2 downward spikes at 21:18:45z and 21:31:09z with amplitude of approximately 700m in the AltGPS.3(Novatel unit) data.  Since the ALTPA data was bad for nearly the first half of the flight, an edited .AXC file was created to filter out the two spikes of erroneous data in AltGPS.3.  AltGPS.3x should be chosen as reference.

The position data for the GPS.2 (Honeywell unit) has many scattered gaps throughout the flight in the position data.  GPS.3x is advised as the best performing GPS at this time.

PSM.1 and PSM.2 differed by 0.34mb upon taxiing at departure and 0.57mb while taxiing upon return indicating neither looked to have wandered throughout the flight.

From 22:24z to 23:15z an example can be seen of a disagreement in our GPS altitude sources.  GPS.3 (Novatel) seems to read about a 20m difference compared to GPS.1 and GPS.1 altitudes.  Still with this disagreement it is advised to use GPS.3x.