## Slide 1: AVAPS User Conference 2018
Contact Info
MSgt Thomas Barnaby
210-842-4532
Thomas.barnaby@us.af.mil

## Slide 2: 2016 SONDE DROPS
TROPICAL – 674
WINTER STORMS – 317 TOTAL = 1269
TRAINING – 278
2017 SONDE DROPS
TROPICAL – 741
WINTER STORMS – 38 TOTAL = 963
TRAINING – 194

## Slide 3
Chart: 2016 Failures
| Terminate Early | Bad Sonde | Bad Battery | Early Launch Detect | Late Launch Detect | No Launch Detect | Fast Fall | Weak Tele/Bad PTU
Series 1 | 12 | 27 | 10 | 9 | 6 | 30 | 25 | 24
143 Total
Failure rate: 11.2%
Chart: 2017 Failures
| Terminate Early | Bad Sonde | Bad Battery | Early Launch Detect | Late Launch Detect | No Launch Detect | Fast Fall | Weak Tele/Bad PTU
Series 1 | 16 | 64 | 14 | 14 | 16 | 38 | 39 | 36
237 Total
Failure rate: 24.6%

## Slide 4
- Bad sondes related to GPS antenna issue were the biggest problem we saw.
- Winds at the bottom of the drop profile during the strongest storms of the year seem to have winds drop off drastically at surface. Waves, bouncing signal, or other cause?
Chart: 2017 Failures
| Terminate Early | Bad Sonde | Bad Battery | Early Launch Detect | Late Launch Detect | No Launch Detect | Fast Fall | Weak Tele/Bad PTU
Series 1 | 16 | 64 | 14 | 14 | 16 | 38 | 39 | 36
237 Total
Failure rate: 24.6%

## Slide 5
Fixed Issues
Sonde GPS antenna should no longer be an issue.
Funding for sondes this year should not be an issue.
ASPEN/AVAPS versions
AVAPS2 3.9.5
ASPEN 3.3-304

## Slide 6
New Business
Windows 10 was added to all drop computers during testing of new NRD41s; issues arose with computers crashing. On-going troubleshooting.
Aircraft going through approval for AXBT receiver equipment to be permanently installed. Possible requirement added to NHOP during IHC conference. Also, another launcher called the Derringer may be used instead to speed up approval.
Transfer process for drop files in-progress, may use google drive for transferring drop data in the interim? Permanent solution needed.

## Slide 7
Contact Info
MSgt Thomas Barnaby
210-842-4532
Thomas.barnaby@us.af.mil