## Slide 1: How do we get good data from Dropsondes?
And what the heck does ASPEN do anyway???????
Kathryn Sellwood UM/CIMAS
Updated June 1, 2022

## Slide 2
IN THE AIR
DROPSONDE MEASURES:
Pressure, Temperature and Relative Humidity twice per second
Wind Speed 4 times per seconds
AVAPS INTERFACE:
- Pre-launch data starts recording when sonde is initialized (cabin)
Launch time set when the drop exits the aircraft & chute is released
Flight level Pressure, Temperature, Humidity, Altitude and Location pulled from
the flight-level data stream and inserted after the launch line
GPS satellite and aircraft motion are removed from drop motion
Wind speed and direction are retrieved
Measurements written to “D” file
Sonde ending (splash) is set by operator once sonde stops transmitting

## Slide 3
AVAPS-D02 P00 120915211 150821 161928.75 817.01 22.71 31.04 223.39 162.32 5.23 -48.348649 14.112884 AVAPS-D02 P10 120915211 150821 161929.00 9999.00 99.00 999.00 223.46 161.71 4.96 999.000000 99.000000 AVAPS-D02 P11 120915211 150821 161929.25 819.46 22.53 999.00 999.00 999.00 99.00 999.000000 99.000000 AVAPS-D02 P11 120915211 150821 161929.50 9999.00 99.00 999.00 999.00 999.00 99.00 999.000000 99.000000 AVAPS-D02 P11 120915211 150821 161929.75 9999.00 99.00 999.00 999.00 999.00 99.00 999.000000 99.000000 AVAPS-D02 P11 120915211 150821 161930.00 9999.00 99.00 999.00 999.00 999.00 99.00 999.000000 99.000000 AVAPS-T02 LAU 120915211 150821 161929.75
AVAPS-D02 A00 120915211 150821 161929.00 660.91 15.08 69.80 230.89 21.95 6.65 -48.347000 14.115000 AVAPS-D02 S00 120915211 150821 161930.25 659.28 21.43 33.62 223.69 154.93 2.66 -48.347116 14.114466 AVAPS-D02 S10 120915211 150821 161930.50 9999.00 99.00 999.00 224.65 146.75 0.09 999.000000 99.000000 AVAPS-D02 S00 120915211 150821 161930.75 660.68 20.27 37.03 224.52 142.40 -1.69 -48.346647 14.114931
P = Pre-Launch
LAU = Launch
A = Flight Level
S = Sounding
Raw D-files:
DYYYYMMDD_HHMMSS.X
DYYYYMMDD_HHMMSS_P.X
DYYYYMMDD_HHMMSS.X.FLIGHTID
Files with “P” in suffix have pre-launch data removed
X=channel # (1-8)
BEFORE ASPEN

## Slide 4
Dropsonde Parachute is designed the keep
the instrument in the optimal vertical
Alignment and control fall speed.
Fall Speed:
Assumption of ~constant fall speed is critical
-to retrieve wind speed from GPS measured
motion
-to assign correct heights and times to
measurements
-to determine velocity pressure
correction
Orientation:
-Controls air flow over sensors
-Protects sensor from elements and solar
radiation
-Determines dynamic response to wind
-Determines dynamic effects on temperature
and pressure
Notes: Measurements are not instantaneous

## Slide 5
IN ASPEN – Quality Control to Remove Questionable Data
- Flag data when communications fail (P/G)
Top of sounding data removed during adjustment to environment (Ae)
Limit check – remove data outside of valid range (Lc)
RH < 0.1% set to 0.1% (after dynamic T correction)
RH > 100% set to 100%
Remove winds when fewer than 6 GPS satellites are present (Sc) – previously 4
Remove winds when vertical velocity differs from hydrostatic estimate by > 2.5 m/s
(after vertical velocity calculation) (Vc)
- Remove winds with error estimates > 1.5 m/s (We) below 100m
Remove winds with error estimates > 1.0 m/s (We) below 10km altitude and above 100m
Remove winds with error estimates 1.5 m/s > (We) above 10km altitude
Compare data with nearby points “Buddy Check” based on maximum slope (Bc)
PTH “Outlier Check” based on standard deviation of the data from a linear fit (Oc)
Remove data when it deviates too much from a filtered profile (Fc)

## Slide 6
IN ASPEN - Dynamic Corrections to Measurements
Pressure data is corrected to account for the dynamic effect of increased air flow velocity over the pressure sensor port assuming constant fall speed
P = P(measured ) – ρv2
- Hypothetical fall speed calculated from sonde mass, parachute size, drag coefficient and
PTRH measurements assuming hydrostatic balance
W= (2Mg/CDρA) 1/2
Wind data is corrected to account for lag in response to wind shear (Laly and Leviton 1958)
Wind speed and direction converted into u v components
Temperature data is corrected to account for time lag between measurement and reaching the receiver (after pressure smoothing is applied)
T = T(measured) + δT/δz *δz/δt*τ
Hypsometric equation integrated upwards from the surface to calculate geopotential height using measured PTRH ( final step)
z1-z2= Rd Tv / g * ln(P1/P2)

## Slide 7
IN ASPEN – Final Filtering to Smooth the Profile
10 second filter applied to Pressure, Temperature, Relative Humidity and Horizontal Wind
Vertical Velocity is computed
5 second filter applied to Vertical Velocity
Create surface observations
Create output files
BEFORE AFTER

## Slide 8
Goal of Dropsonde Scientist:
Provide additional layer of quality control
Inform flight director of significant measurements
Ensure timely transmission of dropsonde data and prevent bad data from
entering the data stream
ASPEN VERSION FOR 2022: 3.4.7
Download ASPEN:
https://www.eol.ucar.edu/software/aspen
Access to User Manual and detailed information about QC and formulas:
https://ncar.github.io/aspendocs/index.html
Changes to ASPEN for 2022
Prevents ASPEN from crashing if flight level data is removed
Pairs latitude and longitude for keep / ignore attributes
Fixes incorrect labeling
Allows comments to be added to Netcdf attributes
Level creation is now consistent across operating systems