Flight ID: 070912I

Sensor or system						Number or Name

Inertial								ine2

Accelerometer							acci2

Temperature Probe						tt2

Dew Point Probe						tdm2X (Vigilant)

Altitude (for vertical wind)					alti1 (GPS aided inertial alt)

Static Pressure							psf (Rosemount Fuselage)

Dynamic Pressure						pqf1 (Rosemount Fuselage)

Time Source							Micro 99

Dropsondes Deployed 						Good:0  	Bad:0  	(# not sent:0 )

Local Met Data:		Takeoff (1422 Z)		Landing (1943 Z)

Aircraft Static Pressure		 1016.6	 mb			1005.2 mb

Tower Pressure			1016.7	 mb			1005.0 mb

Notes

Data Gaps: 142651Z-142656Z   143401Z-143409Z   153051Z-154440Z   170451Z-171430Z

There are small fluctuations in TT2 at 191615Z-191636Z. I decided to leave these in because it appears that they might be natural fluctuations due to atmospheric turbulence as indicated by similar small fluctuations in roll and dynamic pressures.

TDM2 went through a balancing phase between 194450Z-145848Z causing high dewpoint temperatures. This data was replaced with data from TDM1 and saved into the variable TDM2X.

A left 60º roll maneuver was performed at around 193418Z-193515Z

Aircraft INE positions were re-navigated with respect to GPS

Radar altitude from the HG9550 radar altimeter is in the RA-232 variable slot. GPS aided INE1 altitude is in the RA-159 variable slot. The RA-232 and RA-159 altimeters are no longer installed on the aircraft.

The HG9550 radar altimeter data has a consistent 1% positive bias above 1,500 feet that was discovered after flight. The correction formula is:

(((reported altitude -  1,500) * 0.988) +1,500) = corrected altitude

This formula should only be needed for this hurricane season’s flight data as the problem should be corrected automatically during next season’s flights.

There is no King Liquid Water Content (KLWC) sensor installed.

There is no Maycomm Water Vapor Sensing System (WVSSII) installed

There is no TT3 installed.

All other instruments appeared to work optimally.