Flight ID: 071102I

Sensor or system						Variable Name

Inertial								ine1

Accelerometer							acci1X

Temperature Probe						tt1

Dew Point Probe						tdm2 (Vigilant)

Altitude (for vertical wind)					ra159 (GPS aided inertial alt)

Static Pressure							psf (Rosemount Fuselage)

Dynamic Pressure						pqf1 (Rosemount Fuselage)

Time Source							Micro 99

Dropsondes Deployed 						Good: 21  	Bad:1	     AXBT: 16

Local Met Data:		Takeoff (0008Z – 03 Nov)	Landing (0818Z – 03 Nov)

Aircraft Static Pressure		 1012.8	 mb			1015.4	 mb

Tower Pressure			1013.3	 mb			1015.4	 mb

Notes

Data Gap: 001541z-001547z

There were spikes in the accelerometer data at: 015237z-015240z and 045237z-045240z. These spikes were removed from the acci1 variable and the data was saved into the acci1X variable for meteorological calculations.

There were periods where the dewpoint temperature exceeded the ambient temperature resulting in RH values > 100%. This was likely due to heavy rain, a wet-bulb effect on the total temperature sensor, and/or an artificial warming of the dewpoint sensor as it tried to burn off excess moisture. Corrections were not made to the data.

Aircraft INE positions were re-navigated with respect to GPS

Radar altitude from the HG9550 radar altimeter is in the RA-232 variable slot. GPS aided INE1 altitude is in the RA-159 variable slot. The RA-232 and RA-159 altimeters are no longer installed on the aircraft.

The HG9550 radar altimeter data has a consistent 1% positive bias above 1,500 feet that was discovered after flight. The correction formula is:

(((reported altitude -  1,500) * 0.988) +1,500) = corrected altitude

This formula should only be needed for this hurricane season’s flight data as the problem should be corrected automatically during next season’s flights.

There is no King Liquid Water Content (KLWC) sensor installed.

There is no Maycomm Water Vapor Sensing System (WVSSII) installed

There is no TT3 installed.

All other instruments appeared to work optimally.