## Sheet1
PNE 2018 (March 7 - April 14)
Charles Featherstone
Titrator | #3
Burret | #1
Thio | 17.5g/L
Bottles | Green
wheaton | #1
BLANKS 03/08/2018
FINAL BLANK
R1 | R2 | final | R1 | R2
76 | 64.78 | 11.22 | 65.53 | 64.78 | 0.75
72.76 | 65.28 | 7.48
76 | 66.28 | 9.72
70.52 | 66.25 | 4.27
72.51 | 66.03 | 6.48
71.51 | 65.28 | 6.23
65.53 | 64.78 | 0.75
| Average | 5.8217
STANDARDS 03/08/2018
new bottle
Standard | Endpoint
1 | 724.17 | to high did not use
2 | 718.63 | to high did not use
3 | 715.83
4 | 714.38
5 | 714.38
Average | 714.8633
STDEV | 0.83716
Cleaned electrode with nitric acid |  |  |  | Tightened dispensers
Dummy samples run after cleaning with nitric acid and before doing
STANDARDS 03/11/2018 |  |  |  | a new standard
Changed slope to 4 and speed to 0.8 |  |  |  | Dummy Samples from UW
Standard | Endpoint |  |  | Dummy 1 | 206.472 | First sample run did not use in average
1 | 720.89 | to high did not use |  | Dummy 2 | 205.051
2 | 714.24 |  |  | Dummy 3 | 204.758
3 | 715.2 |  |  | Dummy 4 | 204.512
4 | 717.86 | to high did not use |  | Dummy 5 | 205.161
5 | 714.05 |  |  | Dummy 6 | 204.65
Average | 714.4967 |  |  | Average | 204.8264
STDEV | 0.61647 |  |  | STDEV | 0.27251
BLANKS 03/24/2018
R1 | R2 | final | FINAL BLANK
72.47 | 66.43 | 6.04 | R1 | R2
69.95 | 65.68 | 4.27 | 70.2 | 69.7 | 0.5
69.95 | 64.42 | 5.53
70.71 | 65.68 | 5.03
64.92 | 65.17 | -0.25
71.97 | 66.18 | 5.79
64.67 | 65.93 | -1.26
72.72 | 66.68 | 6.04
70.96 | 65.42 | 5.54
70.96 | 65.42 | 5.54
70.46 | 65.68 | 4.78
70.2 | 69.7 | 0.5
| Average | 3.9625
STANDARDS 03/24/2018
Standard | Endpoint
1 | 714.26 | Did not use to high
2 | 708.68
3 | 707
4 | 707.36
Average | 707.68
STDEV | 0.88453
NOTES
Start of Station 7 the burette #1 broke and was replaced with burette 41. The first 2 samples were lost.
Station 016 switched to forward winch
Analyzed Flask 116R (sta. 20) after Flask 118 from station 19, then finished running station 19
Ran new standard and blank before Station 21. Refilled Reagent 1 and 2, acid, thio and KIO3
Changed slope to 4.5 starting with station 034
Station 039 may have left flas number as 245, instead of entering 246
Station 043 Flask 193, Niskin 20 was run as Niskin 1
Station 044 Flask 195-198, run date was entered as 29th, but should be changed to the 30th
Station 045 1st sample Flask 235 was entered as Niskin 1, should be Niskin 3
Station 045 2nd sample Flask 236 was entered as Niskin 2, should be Niskin 4
Station 048 last sample Flask 24R was entered as Niskin 2, should be Niskin 22