Some comments on the installation of a thermosalinograph (TSG)

The volume intake of a micro TSG takes very low pressure.  Sometimes, a water constrictor with a flow meter is installed to reduce the flow rate.

The TSG has no housing.

The TSG needs to be in the engine room, easily accessible but outside the engineer’s way.

The main part of the installation is done when the ship is in port, but with the help of the ship engineering.

The rest of the connection needs to be done when the ship is in transit.

All installation requires approximately 8 days of full time work of two technicians.

It is preferable that the TSG is installed not more than 10 feet (25 max, but with pump) from the water intake; otherwise there are problems with pumping water to the TSG.

The interface box may need to be installed close to the TSG, The ship engineering will decide where to take the water outflow.

It would be preferable to have all instruments properly labeled for when students take a tour to where the TSG is placed.

- The data are collected in a computer that is installed several decks above, where the students are.

The GPS antenna is connected to this computer.

The data transmission to NOAA is done through a computer installed in the bridge.

- There will be an antenna to transmit through Iridium

The installation of the temperature sensor (SBE38) is the most complicated part of the process.  It needs to be installed close (1 to 3 feet) from the water intake.  This installation cannot be done when the ship is moving.