Explorer of the Seas METADATA - 2009

Class of Data:  Surface ocean carbon dioxide concentrations

Dataset Identifier:  AOML_Explorer

One File:  Explorer_2009

Statement of how to cite dataset:

Explorer website: http://www.aoml.noaa.gov/ocd/gcc/explorer_introduction.php

These data are made freely available to the public and the scientific community

in the belief that their wide dissemination will lead to greater understanding

and new scientific insights.  The availability of these data does not

constitute publication of the data.  We rely on the ethics and integrity of

the user to assure that AOML receives fair credit for our work.  Please send

manuscripts using this data to AOML for review before they are submitted

for publication so we can insure that the quality and limitations of the data

are accurately represented.

Measurement platform identifier:  Explorer of the Seas

Cruise Information:

Weekly cruises to Bermuda, Canada or through the Caribbean departing

Bayonne, New Jersey.

Project Information:

This project was sponsored by the underway pCO2 on ships project of the

NOAA Climate Program and represents a collaboration between Royal Caribbean

International, the University of Miami's Rosenstiel School of Marine

and Atmospheric Science, the National Oceanic and Atmospheric

Administration, the National Science Foundation, the Office of

Naval Research, and the National Aeronautics and Space Administration.

Scientists responsible for technical quality of dataset:

Rik Wanninkhof

NOAA/AOML/Ocean Chemistry Division

4301 Rickenbacker Causeway

Miami, Florida 33149

Rik.Wanninkhof@noaa.gov

Dr. Denis Pierrot

NOAA/AOML/Ocean Chemistry Division

4301 Rickenbacker Causeway

Miami, Florida 33149

Denis.Pierrot@noaa.gov

Contact person for this dataset:

Betty Huss

NOAA/AOML/Ocean Chemistry Division

4301 Rickenbacker Causeway

Miami, Florida 33149

Betty.Huss@noaa.gov

Timestamp for initial submission of dataset: 9/09/10

Timestamp for the most recent update of dataset:  9/09/10

Timestamp period the dataset refers to: 3/16/2009 – 11/12/2009

Geographic area the dataset refers to:

13 N to 46 N

88 W to 60 W

List of variables included in this dataset:

COLUMN   HEADER                        EXPLANATION

1.	   Group_Ship:	               AOML_Explorer for all Explorer of the Seas data.

2.       Cruise_ID:		         Cruise Name.

3.	   JD_GMT:	                     Decimal year day.

4.	   DATE_UTC_ddmmyyyy:	         UTC Date.

5.	   TIME_UTC_hh:mm:ss:	         UTC time.

6.	   LAT_dec_degree:	         Latitude in decimal degrees (negative values

are in the southern hemisphere).

7.	   LONG_dec_degree:	         Longitude in decimal degrees (negative values

are in the western latitudes).

8.	   xCO2_EQU_ppm:                 Mole fraction of CO2 in the equilibrator

at equilibrator temperature (Teq) in parts

per million.

9.	   xCO2_ATM_ppm:                 Mole fraction of CO2 in air in parts per million.

10       xCO2_ATM_interpolated_ppm:    xCO2 interpolated in parts per million.

11.	   Pres_EQU_hPa:                 Barometric pressure in the lab in hectopascals

(1 hectopascal = 1 millibar).

12.	   Pres_ATM@SSP_hPa:	         Barometric pressure from ship's barometer,

corrected to sea level in hectopascals

(1 hectopascal = 1 millibar).

13.	   TEMP_EQU_C:                   Temperature in equilibrator water in degrees

centigrade.

14.	   SST_C:                        Temperature from the ship's thermosalinograph in

degrees centigrade.

15.	   SAL_permil:                   Salinity from the ship's thermosalinograph on the

Practical Salinity Scale.

16.	   fCO2_SW@SST_uATM:            Fugacity of CO2 in seawater in microatmospheres.

17.      fCO2_ATM_interpolated_uatm:  Fugacity of CO2 in air in microatmospheres.

18.	   dfCO2_uatm:                  Sea water fCO2 - air fCO2 in microatmospheres.

This uses the interpolated air measurements.

19.	   WOCE_QC_FLAG:                Quality control flag for seawater fCO2 values

with 2 = good value and 3 = questionable.

20.      QC_SUBFLAG:                  Description of the problem with the questionable

Data.

Narrative description of system design:

CO2 ANALYTICAL SYSTEM:

The concentration of carbon dioxide (CO2) in surface ocean water is determined

by measuring the concentration of CO2 in gas that is in contact with the water.

Surface water is pumped from an inlet in the ship's bow to the equilibration

chamber.  The equilibration chamber has an enclosed volume of gas, or headspace,

and a pool of seawater that continuously overflows to a drain.  As the water

flows through the chamber, the dissolved gases (like CO2) partition between the

water and the headspace.  At equilibrium, the ratio of CO2 in the water and in

the headspace is influenced most by temperature, and that relationship is known.

By measuring the concentration of CO2 in the headspace and the temperature in

the chamber, the partial pressure (or fugacity) of CO2 in the surface water

can be calculated.

INSTRUMENT DESCRIPTION:

The general principle of instrumental design can be found in Wanninkhof and

Thoning (1993), Ho et al. (1995), Feely et al. (1998) and Pierrot et al.

(2009).  The concentration of CO2 in the headspace gas is measured using

the adsorption of infrared (IR) radiation, which results from changes in

the rotational and vibrational energy state of the CO2 molecule.  The

LI-COR detector passes IR radiation through two 6" cells.  The reference

cell is flushed with a gas of known CO2 concentration.  The sample cell

is flushed with the headspace gas.  A vacuum-sealed, heated filament is

the broadband IR source.  The IR radiation alternates between the two

cells via a chopping shutter disc spinning at 500 Hertz.  An optical filter

selects an adsorption band specific for CO2 (4.26 micron) to reach the

detector.  The solid state (lead selenide) detector is kept at -12

degrees C for excellent stability and low signal noise (less than 0.2 ppm).

Several steps are taken to reduce interferences and to increase the accuracy

of the measurements.  The IR adsorption band for water vapor is close to the

band for CO2.  After the equilibration chamber, the headspace travels through

a drying trap to remove water vapor.  During each analysis, the headspace gas

is compared to a reference gas of zero CO2 concentration.  To improve the

accuracy of the measurements, three different gaseous standards for CO2 are

analyzed every three hours instead of the headspace gas.

Analyzer:  LI-COR 7000 infrared (IR) analyzer.

Method of Analysis:  Differential analyses relative to a reference gas which

is close to the CO2 concentration of the middle standard.  Measures dried

equilibrator headspace gas.  Gas flow is stopped prior to IR readings.

Drying Method:  The equilibrator headspace sample gas first goes through an

air filter and a thermoelectric refrigerator (~6 C).  The sample and standard

gases pass through a Perma Pure (Nafion) dryer and a short column of magnesium

perchlorate before reaching the analyzer.  The counter flow in the Perma Pure

tube is the reference gas.

Equilibrator (setup, size, flows):  The equilibrator was fabricated using a

filter housing (ColeParmer, U-010509-00) with ~0.5 L water reservoir and

~0.8 L gaseous headspace.  Water flow rate is ~1.5 L/min.  Headspace

recirculation rate is ~80 ml/min.

Narrative statement identifying measurement method for each required parameter:

CALCULATIONS:

The mixing ratios of ambient air and equilibrated headspace air are calculated

by fitting a second-order polynomial through the hourly averaged response of

the detector versus mixing ratios of the standards.  Mixing ratios of dried

equilibrated headspace and air are converted to fugacity of CO2 in surface

seawater and water saturated air in order to determine the fCO2.  For ambient

air and equilibrator headspace the fCO2a, or fCO2eq is calculated assuming

100% water vapor content:

fCO2eq = xCO2eq(P-pH2O)exp(B11+2d12)P/RT

where fCO2eq is the fugacity in the equilibrator, pH2O is the water vapor

pressure at the sea surface temperature, P is the atmospheric pressure

(in atm), T is the SST or equilibrator temperature (in K) and R is the

ideal gas constant (82.057 cm^3·atm·deg^-1·mol^-1).  The exponential term

is the fugacity correction where B11 is the second virial coefficient of

pure CO2

B11 = -1636.75 + 12.0408T - 0.032795T^2 + 3.16528E-5 T^3

and

d12 = 57.7 - 0.118 T

is the correction for an air-CO2 mixture in units of cm^3·mol^-1 (Weiss, 1974).

The calculation for the fugacity at SST involves a temperature correction term

for the increase of fCO2 due to heating of the water from passing through the

pump and through 5 cm ID PVC tubing within the ship.  The water in the equilibrator

is typically 0.16 °C warmer than sea surface temperature. The empirical

temperature correction from equilibrator temperature to SST is:

fCO2(SST) = fCO2(eq) Exp (-0.0423 (Teq-SST))

where SST is sea surface temperature and Teq is the equilibrator temperature

in degrees °C.

Sampling Cycle:

The system runs on a cycle during which 4 standard gases, 5 air samples

from the bow of the ship and 50 surface water samples from the

equilibrator head space are analyzed on the following schedule:

- Zero and span of Licor

- Four standard gases

- Five air samples

- Fifty equilibrator headspace gas samples

- Repeat steps 2 – 4 nine more times

- Restart from step 1

NOTES ON DATA:

Columns have a default value of -9 in case of instrument malfunction,

erroneous readings or missing data.  Furthermore, if a suspicious xCO2 value,

pressure or temperature value is encountered, the fCO2 is not calculated.

Analytical Instrument Manufacturer/Model:

The Explorer of the Seas system is built by General Oceanics.  The

analyzer is a LI-COR 7000 infrared analyzer.

Standard Gases and Reference Gas:  The three standard gases come from CMDL

in Boulder and are directly traceable to the WMO scale.  The reference gas

is a zero CO2 gas from a commercial company.  While individual data points

above 460 or below 300 may not be accurate, the general trends should be

indicative of the seawater chemistry.

Description of any additional environmental control:

The system is located in the bow thruster space of the Explorer of the

Seas.

Resolution of measurement:

The resolution of the instrument is better than 0.1 ppm.

Estimated overall uncertainty of measurement:

The xCO2eq measurements are believed accurate to 0.1 ppm.  The

fCO2@SST measurements are believed to be precise to 0.2 ppm.

List of calibration gases used:

The standards used on cruises are:

STANDARD      TANK #       CONCENTRATION        VENDOR

STD1          CA02040         319.82			CMDL

STD2          CA02030         377.14			CMDL

STD3          CA02020         549.89             CMDL

Traceability to an internationally recognized scale (including date/place

of last calibration made):

All standards are obtained from NOAA/CMDL, now called the Global

Monitoring Division of the Earth Research Laboratory and are directly

traceable to WHO scale.

Uncertainty of assigned value of each calibration gas:

The uncertainty based on pre and post cruise calibrations is less

than 0.05 ppm.

Units:

All xCO2 values are reported in parts per million (ppm) and fCO2 values

are reported in microatmospheres (uatm) assuming 100% humidity at the

equilibrator temperature.

Bibliography:

DOE (1994). Handbook of methods for the analysis of the various parameters of the

carbon dioxide system in sea water; version 2. DOE.

Feely, R. A., R. Wanninkhof, H. B. Milburn, C. E. Cosca, M. Stapp and P. P.

Murphy (1998).  A new automated underway system for making high precision

pCO2 measurements onboard research ships. Analytica Chim. Acta 377: 185-191.

Ho, D. T., R. Wanninkhof, J. Masters, R. A. Feely and C. E. Cosca (1997). Measurement

of underway fCO2 in the Eastern Equatorial Pacific on NOAA ships BALDRIGE and

DISCOVERER, NOAA data report ERL AOML-30, 52 pp., NTIS Springfield.

Pierrot, D., C. Neil, K. Sullivan, R. Castle, R. Wanninkhof, H. Lueger, T. Johannson,

A. Olsen, R. A. Feely, and C. E. Cosca (2009).  Recommendations for

autonomous underway pCO2 measuring systems and data reduction routines,

Deep-Sea Res II, 56, 512-522.

Wanninkhof, R. and K. Thoning (1993).  Measurement of fugacity of CO2 in surface water

using continuous and discrete sampling methods. Mar. Chem. 44(2-4): 189-205.

Weiss, R. F. (1974). Carbon dioxide in water and seawater: the solubility of a

non-ideal gas.  Mar. Chem. 2: 203-215.

Takahashi, T., J. Olafsson, J. G. Goddard, D. W. Chipman, and S. C. Sutherland

(1993).  Seasonal variation of CO2 and nutrients in the high-latitude surface

oceans: a comparative study, Global Biogeochem. Cycles, 7, 843-878.

Comments related to the individual legs:

EX0908: A faulty valve was discovered at the end of the cruise preventing

the air to be returned to the equilibrator. The air was then pulled out of

the equilibrator and just vented out after going through the analyzer. The

loss of air was made up by ambient air going through a pre-equilibrator.

It is not known when the valve went faulty. We estimate the uncertainty on

the xCO2 values to be around +/- 5 ppm.  All data starting on March 22 was

obtained with a reduced EQU flow to try to minimize the loss.  An offset of

-3.11 minutes was applied to the SST data (SST was lagging compared to

EQU Temp) prior to Delta T calculations.  Water flow problems seem to have

occured around JDay 81.25 and 81.3 and the xCO2 data shows greater variation

so the data has been flagged 3.

EX0924:  Near the end of the cruise, it was suspected that the three-way valve

that allows analyzed headspace to return to the equilibrator was faulty.  The

analyzed headspace was being vented; the removed headspace was made up with ambient

air inside the wet box.  Around 12:00 on 31 July, the extra outside air exiting the

ballast (source gas for Nafion counterflow) was directed into a vented enclosure.

The ambient gas being pulled into the equilibrators came from this enclosure.

During this cruise the ship’s scientific infrastructure was being converted to

full automation.  Data was not collected during the early and late parts of the

cruise because of control software development issues.  Data was collected during

the southernmost transits using the control software’s earliest version.  The

connections to and configuration of auxiliary sensors (TSG, SST, etc.) were also

being optimized.  These changes resulted in loss of data being logged by the pCO2

software program.

On 26 July between 15:16 and 16:42, there was no GPS position recorded.  Since

this gap occurred during a ship transit day, the positions were extrapolated

from the constant speed and direction between 13:54 and 14:17.

On 26 July between 13:15 and 22:21 the TSG, SST, and atmospheric pressure

sensors were not logged.  For this time interval, the default value of 35 was

assigned to salinity.  The difference between the equilibrator and SBE38

temperatures was calculated for the 1190 analyses when both were logged

(average= 0.17 +/- 0.25 C).  For the above time interval, the SST was

estimated by subtracting 0.15 C from the equilibrator temperature.

The difference between the Licor and atmospheric pressures was calculated for

the 1056 analyses when both were logged (average= 4.96 +/- 0.33 mbar).  When the

atmospheric pressure was not logged, it was estimated by subtracting 4.96 mbar

from the Licor pressure.  This estimation was done during the above time interval,

on 30 July between 17:17 and 19:61, and between 30 July 20:43 and 31 July 08:08.

There were single analyses that did not successfully record all the auxiliary

sensors.  These single drop-outs were interpolated from the neighboring data.

The frequency of the drop-outs decreased significantly as the cruise progressed.

Several times during the transits between islands, the air analyses were slightly

higher than expected.  Terrestrial influence of the atmospheric CO2 concentrations

was suspected, and the results are reported.

The applied time offset between the SST and equilibrator temperature data is 0.3

minutes.  The changes in temperature recorded by the SBE38 were recorded in

the equilibrator 18 seconds later.

The SST and salinity data are provided as a reference.  These data have not

undergone rigorous quality control processes.