# GasEx 98 Leg 2 pCO2 Data Reduction Notes

Cruise start:	May 25, 1998	Lisbon Portugal

Cruise end	June 26, 1998	Punta Del Gado Azores

Standard Tanks:

Low:	275.63	CC114999	1920 PSI

Mid:	297.74	CC117904	800 PSI

Hi:	360.62	CC106641	1990 PSI

Equilibrator Mercury Thermometer:	Brooklyn Serial No. 75293

Thermistor Algorithm:	102.04 – 48.88V + 10.785V2 – 1.2714V3

Thermistor Calibration performed on April 8, 1998.

Notes from Green Log Book:

New reference mass flow meter installed prior to Leg 2.  Serial # G6736, scale 0 – 50.  Original was used for SF6 system.  The program algorithm for converting voltage to flow was not changed, so all flows should be multiplied by 4 to give accurate readings as the original MFM had a scale of 0 – 200.

During this leg WHOI teed off our system at the equilibrator and low standard tank.  Their system continuously measured the water fCO2 using a Licor analyzer and exposed some leakage problems in the system.  The main leaks occurred in the garden hose fittings and were repaired on May 29.  Also, the vent lines from the equilibrator fed into the hydro lab instead of outside and some leakage also occurred in the condensor glassware.  These problems are noted below.

May 26, 2000 – 2300:	WHOI teeing off system.

May 27, 1100 – 1246:	Pressure check of system.  STD 1 regulator accidentally turned off.

May 27, 1530 – 1546:	Moved STD1 regulator.  No reference flow.

May 28, 1840 – 1850:	WHOI working on their YSI.

May 29, 1830 – 2030:	Moved Dekoron vent lines from equilibrator which were venting into the Hydro Lab.  They now vent to open atmosphere above the van aft of the Hydro Lab.  Also closed off secondary drain (water) on the bottom of the equilibrator.  Replaced leaking washers in garden hose fittings.  After these repairs, WHOI no longer observed the spikes that were previously apparent in their data.

May 30, 1500 – 1800:	YSI flow shut off.  From 1400 to 1700 water flows increased because CMDL turned off their equilibrator.

May 30, 2000:	STD1 reference and sample flows found to be low.  Increased them.

May 30, 2330 – 2400:	Equilibrator water level too low.  Positive pressure in equilibrator caused headspace gas to back up into Licor sample cell, resulting in bad readings.

May 31, 2115 – 2300:	Shut system down to install valve on sample output of Licor.  During air and STD phases, Licor now sends sample output to atmosphere.  During water phase sample output goes to equilibrator as before.

June 1, 1100	Discovered leak in condensors.  Tightened fittings, greased O-ring , and tightened clamps.

June 1, 2342	Valve Mistrip.

June 2, 2049 & 2242	Valve Mistrip.

June 3, 1412 – 1825:	Shut down system to recalibrate YSI.

June 3, 2120 – 2150:	Reference gas disconnected while teed off tank.

June 4, 1545:	Valve Mistrip.

June 7, 2232	Valve Mistrip.

June 9, 545 – 625:	Cleaned equilibrator shower head.

June 11, 2333:	Valve Mistrip.

June 12, 1225 – 1312	Craig Neill changed the circuit controlling the Valco actuator, inserting an Opto-coupler and eliminating the mechanical relays that previously controlled it.  This change eliminated the valve mistrips the plagued the system in the past.

June 12, 1430 – 1600:	Recalibrated YSI O2 sensor.

June 17, 1505 – 1930	Recalibrated YSI O2 sensor.

June 21, 1053:	Increased water flow from ~ 9.9 to ~10.2 L/min.

June 23, 1320:	Last CTD completed.  Started steaming for Azores.

June 24, 1600:	WHOI disconnected their equipment.

June 25, 1315 – 1640:	System hung up while I was explaining it to a future operator.  Mea Culpa.

June 26, 0700:	Seawater supply off.  System shut down.  End of Leg 2.

Below are rows where fixes have been made:

- Interpolated GPS time and lat.

  - Interpolated Std mVs.

- Added row.  Interpolated Jday, lat, lon, mV, SCS, & Eq temp.

- Interpolated mV.

- Interpolated mV.

604,18,32 	Interpolated mV.

- Used previous hour’s air xCO2 value for NEWxCO2.

- Mistrip.  Used next value.

- Mistrip.  Interpolated.

  - Mistrip.  Changed mV to –9.

- Interpolated mV.

  - Interpolated YSI sal.

  - Mistrip.  Used next value for STDs 1 & 2, interpolated STD 3.

  - System down for installation of solenoid valve.  Added rows.  Interpolated Jday, lat, lon, mV, SCS, & Eq temp.

2134,53,62	Interpolated SCS sal.

2225,31,33	Interpolated SCS sal.

2254,67	Interpolated SCS sal.

- Interpolated SCS sal.

2400,29,48	Interpolated SCS sal.

2467,76,88	Interpolated SCS sal.

- Interpolated SCS sal.

2501,11,16,

2527,32	Interpolated SCS sal.

- Interpolated YSI sal & DO.

  - Mistrip.  Used next value for STDs 1 & 2, interpolated STD 3.

  - Mistrip.  Used next value for STDs 1 & 2, interpolated STD 3.

3015-74	Recalibrating YSI – system down.  Added rows.  Interpolated Jday, lat, lon, mV, SCS, & Eq temp.

  - Mistrip.  Used next value for STDs 1 & 2, interpolated STD 3.

- Interpolated SCS sal & temp.

- Mistrip.  Interpolated STD 3 mV.

- Interpolated SCS sal & temp.

  - Mistrip.  Used next value for STDs 1 & 2, interpolated STD 3.

  - Mistrip.  Used next value for STDs 1 & 2, interpolated STD 3.

- Mistrip.  Interpolated STD 3 mV.

  - Mistrip.  Used next value for STDs 1 & 2, interpolated STD 3, changed others to –9.

- Interpolated SCS sal.

  - Mistrip.  Used next value for STDs 1 & 2, interpolated STD 3.

  - Mistrip.  Used next value for STDs 1& 2, interpolated STD 3, changed others to –9.

  - Recalibrating YSI – system down.  Added rows.  Interpolated Jday, lat, lon, mV, SCS, & Eq temp.

- Interpolated pressure.

  - Recalibrating YSI.  All YSI readings no good.

  - Recalibrating YSI - system down.  Added rows.  Interpolated Jday, lat, lon, mV, SCS, & Eq temp.

- Bad flow.  Interpolated STD 3 mV.

10392	Interpolated YSI readings.

10394-440	System hung up.  Added rows.  Interpolated Jday, lat, lon, mV, SCS, Eq Temp, and YSI readings.