Cruise:  ABACO, Eastern Boundary Current

Ship:  R/V Ronald H. Brown

Dates:  March 9- 28, 2006

Chief Scientists:  Molly Baringer,  Lisa Beal

Equipment:  24 bottle CTD with dual C, T, and O2 sensors and LADCP

Number of Stations:  72

Sample Collection

O2:

Shipboard, all stations, all depths

PI:  Molly Baringer, AOML

Analyzed By:  Robert Roddy, AOML

O2 in umol/kg

Halocarbons:

Shipboard

PI:  Dr. Shari-Yvon Lewis, TAMU

Analyzed by:

Nutrients:

23 stations, 391 samples, frozen and returned to lab. Note: in the lab the samples thawed and were analyzed XX weeks after samples were thawed.

PI:  Dr. Jia-Zhong Zhang, AOML

Analyzed By:  Charlie Fischer, AOML

DOC/TOC:

33 stations, 506 samples, samples frozen and shipped to RSMAS.

PI:  Dr. Dennis Hansell, RSMAS

Analyzed by:  Wenhao Chen, RSMAS

Charles Farmer, RSMAS

DIC:

6 stations, 80 samples each 500-ml, 6 sets of duplicate samples

PI:  Dr. Rik Wanninkhof, AOML

Analyzed by:  Esa Peltola, AOML

TALK:

2 stations, 40 samples, 3 duplicates

PI:  Dr. Frank Millero, RSMAS

Analyzed by:  Fen Huang, RSMAS

Sample Analysis

DIC:

Analysis Date: April 4-6, 2006.

Coulometers used: AOML1 and AOML2

Blank range: 15.6-29.2 counts/min

CRM # used and assigned value (include both DIC and salinity): Batch 69, c:1907.63 umol/kg, S:31.569.

CRM value measured:  AOML 1: offset 9.9 umol/kg (1897.7 umol/kg)

AOML-2: offset 7.4 umol/kg (1900.2 umol/kg)

The average run time was 13 minutes, the minimum run time was 9 minutes and the maximum run time was 20 minutes.

Reproducibility: (# samples and average difference): 6 sets of duplicate samples, average difference 1.1 umol/kg.

CRM, salinity and HgCl2 correction applied: yes

Remarks-

A density correction was applied for all the samples.  Due to added HgC12 (measured DIC*1.00037), a volume correction was applied for all the samples.  The average CRM correction was 8.7 umol/kg and was run at the beginning and end of the cell.  If the cell’s lifetime reached it, a CRM was also run at 12 mg of C.  The first CRM of each cell was used for a CRM correction.  There was good agreement between the duplicate samples.  The CTD temperature and salinity was interpolated by Carlos Fonseca for the following samples:

Station 57/Niskin Bottle 2

Station 62/Niskin Bottle 16

The DIC sample bottle number 104 was not found in the sampling sheets.

It was estimated to be station 20/Niskin bottle 8 by Carlos Fonseca and added as such.

During the cruise, there was some confusion about how to secure the top to the sample bottles with rubber bands.  The coulometers had not been used for several months.  Acid delivery valve #8 and drain valve #9 malfunctioned and needed to be replaced.  Forty samples were sent to Dr. Millero's lab for alkalinity analysis.

Comments

The Sample_ID is the rosette niskin bottle number from which the discrete samples were drawn.