Calibration for fluorometer 1990:

Spectrophotometer readings

wavelength               ext. coeff.

750                      0.0054

664                      0.0521

663                      0.0517

647                      0.0513

630                      0.0177

Concentration of chla for live algal culture = 11.85(.0521-.0054) - 1.54(.0513-

.0054) - 0.08(.0177-.0054)

= 0.481725 ug/ml

Therefore, stock solution is 0.482 ug/ml

Diluted stock to 60:40 mixture of 12ml stock: 8ml DMSO

Ranges of concentrations used for calibration:

0.0289

0.0202

0.00867

0.00289

0.00231

0.00173

0.00116

0.000578

0.000145

1/Kc = concentration chla/Fb

scale           mean 1/Kc

31.6x           0.000015

10x             0.000053

3.16x           0.00017

1x              0.000053

Calibrated only on 1x door because all readings were on this door

Tau = Fb/Fa

scale           mean Tau    Tau Factor ( Tau/Tau-1 )

31.6x           1.89        2.12

10x             2.05        1.95

3.16x           2.12        1.89

1x              2.24        1.81

Equation to determine chla conc. from samples :

Conc.( ug/L ) = (1/Kc)*(Tau/Tau-1)*(Fb-Fa)*(vol extr.ml/Vol filt.L)

Equation to determine phaeopigments :

Conc.( ug/L ) = (1/Kc)*(Tau/Tau-1)*(Tau*Fa-Fb)*(vol extr.ml/Vol filt.L)

Calibration 1991:

Spectrophotometer readings

wavelength         ext. coeff.

665                1.771

750                0.325

after acid addition:

665                1.118

750                0.327

Corrected 665 values = 1.446

after acid = 0.791

Chlorophyll a concentration = 26.7(corr. 665 before acid - corr.665 after

acid) = 17.488 ug/ml

Range of concentrations used for calibration:

0.0525 - 0.0001749 ug/ml

1/Kc values:

door      scale         mean 1/Kc     Tau

100x      3.16x         0.0000131     1.78

100x      1x            0.0000481     1.74

1x        10x           0.0006263     2.17

1x        31.6x         0.0001955     2.02

Almost all readings corresponded to these scales. However, there were some

readings on other scales that were not directly calibrated. These

concentrations were calculated using values extrapolated from the values

obtained from the equation :

1/Kc*Tau/Tau-1

The full range of values for 1/Kc*Tau/Tau-1 is:

door       scale         value

100x       1x            .0001131

100x       3.16x         .00003

100x       10x           .0000095

100x       31.6x         .000003

1x         1x            .0116

1x         3.16x         .00367

1x         10x           .00116

1x         31.6x         .000387

Concentrations of chlorophyll a and phaeopigments were calculated using the

same equations as in the 1990 data.

References for all Methods:

Parsons, T.R., Maita, Y. and Lalli, C.M. ( 1984 ) A manual of chemical and

biological methods for seawater analysis. Pergamon press.

Shoaf, W.T. and Lium, B.W. (1976) Improved extraction of chlorophyll a

and b from algae using dimethyl sulfoxide. Limnol. Oceanogr. 21, 926-928

Smith, R.C., Baker, K.S., and Dustan, P. ( 1981 ) Fluorometric techniques

for measurements of oceanic chlorophyll in support of remote sensing. SIO

Ref. 81-17, 14pp., Scripps Inst. of Oceanography, La Jolla, Ca.

#####################################