Equipment Failure and Damage Report

AX081210 High Density XBT Cruise

MK21 USB Failure

The primary MK21 USB unit (S/N 0929) failed inexplicably on 12/09. There was no indication as to the cause of the failure. The unit could not connect to either the PC or the auto-launcher. There was no evidence of the clicking sound which indicates that the unit is functioning. The MK21 unit, as well as all other equipment, was connected to a UPS providing both surge protection and battery backup. The unit was replaced with the secondary MK21 for the remainder of the cruise.

Damage to Auto-launcher mounting

The auto-launcher was displaced from its mounting during extremely poor weather encountered on 12/20. The hose-clamps securing the mounting pole to the railing snapped and the G-clamps used for additional support were loosened, resulting in the auto-launcher falling onto the deck. The auto-launcher mounting frame was buckled either prior to being knocked off, or with the impact of falling onto the deck.

During this period the spray cover for the 8-tube auto-launcher was blown off by the hurricane-strength winds, and was lost into the sea.

Hand-launcher

The hand-launcher was used for the final two drops of the cruise. The penultimate drop registered a constant temperature throughout the profile. On the final drop there was no splash detected when the drop was made. Due to a lack of probes it was not possible to retry these drops and determine if there was in fact a fault with the hand-launcher. It is assumed that the hand-launcher is faulty, but not confirmed.